LIMS REQUIREMENTS SPECIFICATION

Laboratory Information Management System

Document No: LIMS-RS-001  |  Rev 1.2  |  Status: Approved for Vendor RFP  |  Date: 20-Jun-2024

# 1. Purpose & Scope

This document defines the functional and non-functional requirements for a Laboratory Information Management System (LIMS) for the Lube Oil Blending Plant QC laboratory. The LIMS will replace the current Excel-based QC tracking system and integrate with SAP QM, SCADA, and laboratory instruments.

Lab Location: Lube Oil Blending Plant, Quality Control Laboratory (Building 3)

Lab Type: Lubricant blending QC — routine finished product testing, incoming raw material testing, retained sample management

Sample Volume: Approx. 25–40 samples per day (in-process + finished product + raw materials)

Users: QC Analysts (6), QC Manager (1), Shift Supervisors (3 — view only), Commercial team (CoA access only)

Integration: SAP QM, SAP PP-PI, SCADA Historian, Lab Instruments (5 types), Customer CoA portal

# 2. Functional Requirements

Priority: Must Have = mandatory for go-live. Should Have = required within 6 months. Could Have = future roadmap.

| FR-ID | Feature | Requirement Description | Priority | Integration |
| --- | --- | --- | --- | --- |
| FR-001 | Sample Registration | System shall auto-generate unique sample IDs linked to SAP batch number when a batch process order is confirmed. | Must Have | SAP PP-PI |
| FR-002 | Barcode Printing | System shall print barcode labels with sample ID, batch number, product grade, and test list at sample collection point. | Must Have | Label printer |
| FR-003 | Test Scheduling | System shall auto-assign test list to each sample based on product grade rules configured in the LIMS method library. | Must Have | None |
| FR-004 | Instrument Integration | System shall receive raw results directly from: viscometer, COC flash tester, autotitrator, ICP-OES spectrometer, and Brookfield viscometer. | Must Have | Instrument RS-232/Ethernet |
| FR-005 | Specification Checking | System shall automatically compare all results against product-grade specifications and flag out-of-spec results with red highlight and email alert. | Must Have | None |
| FR-006 | SAP QM Integration | System shall send usage decision (Accept / Reject) to SAP QM inspection lot via REST API within 5 minutes of final QC approval. | Must Have | SAP QM |
| FR-007 | CoA Generation | System shall auto-generate Certificate of Analysis in PDF format from approved results, batch header, and customer details. | Must Have | SAP SD |
| FR-008 | Audit Trail | System shall maintain a full audit trail of all result entries, modifications, approvals, and rejections with timestamp and user ID. | Must Have | None |
| FR-009 | Stability Study Management | System shall support periodic re-testing schedules for retained samples over 12, 24, and 36-month periods. | Should Have | None |
| FR-010 | Multi-site Reporting | System shall consolidate QC data across the blending plant and the storage terminal for unified dashboards. | Could Have | None |
| FR-011 | LIMS Mobile Access | System shall provide a mobile-responsive interface for operators to log sampling events and view results from tablets in the plant. | Should Have | Mobile device MDM |
| FR-012 | Trend Analysis | System shall provide trend charts for each test parameter per product grade over user-selectable time ranges. | Should Have | None |

# 3. Non-Functional Requirements

| Category | Requirement |
| --- | --- |
| Performance | Page load time <2 seconds. Result import from instruments <10 seconds. CoA generation <30 seconds for standard grades. |
| Availability | 99.5% uptime during working hours (06:00–22:00). Scheduled maintenance window: Sunday 00:00–04:00. |
| Security | Role-based access control (RBAC). Full audit trail per 21 CFR Part 11 principles. SSO via Active Directory. Data encrypted at rest (AES-256) and in transit (TLS 1.3). |
| Scalability | Support up to 500 samples/day without performance degradation. Extensible to additional lab sites without core system changes. |
| Data Retention | QC results: 10 years. Audit logs: 7 years. CoA records: lifetime of product. Retained sample register: 5 years post-expiry. |
| Disaster Recovery | RTO: 4 hours. RPO: 1 hour. Daily automated backup to separate server. Monthly DR drill. |
| Validation | LIMS must support IQ/OQ/PQ validation. Vendor to supply validation master plan template and test scripts. |
| Training | Vendor to provide: admin training (2 days), user training (1 day per role), train-the-trainer package, and video e-learning modules. |

# 4. Instrument Integration Specifications

| Instrument | Model | Interface | Data Points to Import |
| --- | --- | --- | --- |
| Automated Viscometer | Anton Paar SVM 3001 | Ethernet / LIMS Middleware | KV @40°C, KV @100°C, Density @15°C, VI (calculated) |
| Cleveland Open Cup | Koehler K16300 | RS-232 → USB | Flash Point (°C), Fire Point (°C), Barometric Pressure |
| Autotitrator | Metrohm 916 Ti-Touch | Ethernet / LIMS Middleware | TBN (mg KOH/g), TAN (mg KOH/g), titrant volume, sample weight |
| ICP-OES Spectrometer | Spectro Arcos II | Ethernet / CSV import | Zn, P, Ca, Mg, S, Fe, Cu, Pb, Si (ppm) |
| Brookfield Viscometer | Brookfield DV3T | RS-232 → USB | Viscosity (cP), Temperature (°C), Spindle ID, Speed (RPM) |

# 5. SAP QM Integration Data Flow

| Event | Data Flow & Timing |
| --- | --- |
| Process Order Created (SAP → LIMS) | SAP PP-PI process order creation triggers LIMS via REST API. LIMS creates inspection lot, assigns sample IDs, and queues test list. Latency target: <60 seconds. |
| Sample Results Available (LIMS internal) | Instrument results imported, auto-compared to spec. QC Manager notified for review and approval via LIMS dashboard and email. |
| Usage Decision (LIMS → SAP QM) | On QC Manager approval in LIMS, usage decision (Accept/Reject) posted to SAP QM inspection lot via SAP OData API. Triggers batch stock posting in SAP. Latency: <5 minutes. |
| CoA Generation (LIMS → SAP DMS) | Approved CoA PDF archived in SAP DMS linked to batch. Copy sent to customer portal. SAP billing document can reference CoA number. |